SPECIFICATION

Manufacturer: 			Safelab Systems Ltd

Address:	Airone Building, 8 Beaufighter Road, Weston-Super-Mare, BS24 8EE

Telephone:	+44 (0)1934 421340

E-mail:				enquiries@safelab.co.uk

Website: 			www.safelab.co.uk

Supplier Contact:			Adrian Wyles – email: adrian@safelab.co.uk tel: +44 (0)7398 288292

Product description:		1000mm wide Fixed Ducted fume cupboard

Range:				Airone X2 Fixed Ducted Constant Air volume fume cupboard

Product Code:			Airone 1000X2 Preparation Room

Application: 	For use whenever the operator and the environment need to be protected from gaseous or particulate contaminants. Designed for use in school laboratories and preparation rooms.

Standards:	Designed to comply with CLEAPSS G9 Fume cupboards in Schools. Previously successfully tested by CLEAPSS against the specification in Building Bulletin 88, Fume cupboards in Schools which has been replaced with CLEAPSS G9 Fume cupboards in Schools.

Guarantee:	5 years warranty when Safelab carry out the annual service.

Installation:			Freestanding without the need for fixing to the walls or floor

External dimensions:		1000mm wide x 760mm deep x 2200mm high

Internal dimensions:		950mm wide x 550mm deep x 780mm high

Average inflow:			>0.45 m/s

Air Volume:			600m³/hour

External carcase: 	Manufactured from 16swg mild steel and profiled aluminium aerofoil extrusions and painted in epoxy powder coating

Glazed side panels:		6mm Laminated security glass

Baffle:				6mm solid grade laminate

Lower Aerofoil:			Profiled Stainless Steel

Worktop:			Fabricated white polypropylene spillage tray with drip cup

Sash:				6mm Laminated security glass

Sash operation:	Sliding with counterweights – conforms to BS EN14175-2:2003 as indicated in item 2.1.3 of CLEAPSS G9 Fume Cupboards in Schools document, the sash cannot fall if one of the suspending cable-cords is cut

Air Bypass:	A perforated area on the upper fascia provides bypass air entry, decreasing as the sash is raised. This operates over the full height of the working sash opening providing constant air volume.

Minimum working aperture:	200mm high

Maximum working aperture:	400mm high

Support stand:			Fully welded tubular steel.

Colour:				Light Grey RAL-7035 and White RAL-9016

Storage cabinet (vented):	Manufactured from mild steel and painted in epoxy powder coating

Low Airflow alarm:	Monitors safe airflow at the working apperture

Weight:				84kg excluding stand

Services:

Gas service:			Single gas outlet and fascia mounted control valve

Gas connection:			¼” x 10mm compression fitting

Water tap:			Single cold-water outlet with serrated nozzle and fascia mounted control valve

Water tap connection:		¼” x 10mm push fit adaptor

Drip cup:			Vulcathene Drip cup

Waste:				38mm Vulcathene Bottle trap

Electrical sockets:			Fascia mounted Double 13Amp outlet

Lighting:				LED

Supply:				230V/50Hz, 13Amp